BCD Series controller instructions:
Before turning on system be sure tank is filled 3 inches below top.

System ON/OFF press temperature knob.
The display will indicate two sets of information, left side displays heat status	and right side ultrasonic timer.

Heat upper digits indicates off or set temperature,	lower numbers are the actual	temperature of the solution.

Ultrasonic timer (right side of display) upper digits will start on one (01)	indicating one min. on the timer. Lower numbers will say zero (00) not running.

Temperature settings:
Turn knob clockwise at the left of display to set desire temperature (upper digits)	and the heating system will activate. Maximum temperature is 160°f. Solution	temperature is displayed below set value.

Ultrasonic timer:
Timer preset value is adjusted by turning knob on the right of the display.	Clockwise rotation will increase value. Once desired amount of time is obtained	push knob to start ultrasonics. Lower numbers will count down remaining time.	Press knob to stop ultrasonic. The maximum timed cycle is 98 min.

Continuous operation of ultrasonic is achieved by going past 98 minutes, at that point the display will indicates “ON”. Press the knob to activate ultrasonics.

ULTRASONIC INTENSITY: This knob control varies power output of the generator (ultrasonics) between a minimum of 35watts and the generators maximum rated power output, as the control is rotated CW.

DEGASSING CONTROL OPTION: The left-hand control is used for degassing. During this operation, the generator output is switched off periodically to help allow dissolved gasses to escape. For this mode, the generator has a fixed on period of 110 milliseconds (MS). CW rotation of the control varies the off period between 20 and 70 MS. Full CCW rotation disables the interrupt feature and the generator operates continuously.

Dependable ULTRASONIC equipment for precision cleaning

Made in USA Since 1966